## Slide 1: The Frequency shift measurements of mid-T treated cavities
Rezvan Ghanbari
on behalf of SRF R&D team
14 December 2022

## Slide 2: Too much noise prevent resolution of the dip
In first attempts, low SNR above 9K prevented dip behavior studies
Warm-up / cool-down process to fast to improve SNR by simple VNA adjustments
VNA cannot sample fast enough

## Slide 3: Accurate f0 due to Lorentz fit
Record spectrum of each frequency sweep
Fit the Lorentz curve to each spectrum
Significant improvement above 9K
Below 9K results are identical due to high SNR

## Slide 4: Dip behavior is now resolved

## Slide 5
f vs. T of 5 Cavities
Cavity | Δf / kHz | Δfdip / kHz | ΔTdip / K
1DE19 | 11.45 | 1.10 | 0.04
1AC2 | 11.47 | 1.34 | 0.08
1AC3 | 12.84 | 0.88 | 0.09
1DE09 | 17.69 | 1.35 | 0.23
1RI02 | 21.77 | 1.48 | 0.29
Note: A change of f is observed above TC of Nb.
Temperature of cavity & sensor maybe not in equilibrium due to too fast cool down
Hence a relative definition of the dip position ΔTdip
ΔTdip
Δfdip
Δf

## Slide 6: Summary
Improved resolution of f vs. T measurements with noise reduction procedure
Tested five mid-T treated cavities:
  20h@250°C shows larger f shift as well as larger dip depth
  LG cavity shows lower dip depth and higher f shift compare to FG cavity after the same treatment
  Similar results after the same mid-T treatment in the same furnace
Temperature sensor and cavity might not be in equilibrium due to warm-up rate
Next steps:
Increase statistics: continue frequency shift measurements for mid-T treated cavities
Develop an optimal length of input antenna suitable for all RF measurements
Optimize VNA settings to further reduce noise floor
Improve heating control system in cryostat during cool-down and warm-up

## Slide 7: Thank you
Rezvan Ghanbari
Institute of Experimental Physics
University of Hamburg
E-mail: rezvan.ghanbari@desy.de Phone: (+49) 040-8998-4321
Contact
Deutsches Elektronen-
Synchrotron DESY
www.desy.de