## Slide 1
Joffray GUILLORY, Daniel TRUONG, Jean-Pierre WALLERAND
Conservatoire national des arts et métiers
Laboratoire commun de métrologie LNE-Cnam
IWAA | International Workshop on Accelerator Alignment
October 5-9 2026 at KEK, Tsukuba, Japan
1
A hybrid ADM–IFM system for air refractive index compensation : Target uncertainty of 5 µm (k=1) up to 100 m
French University
National Metrology Institute

## Slide 2
  Laser trackers, Total stations: they operate at a single wavelength.
Issue: the measured distance depends strongly on the atmospheric parameters
Principle of air refractive index compensation
2
  partial pressure of water
At two wavelengths: compensation of the air refractive index with temperature and pressure
  Moist air
  Dry air
The factor A is a constant
A = 21 for 442 nm / 1633 nm wavelengths (group index).
A = 48 for 780 nm / 1560 nm wavelengths (group index).
A = 65 for 532 nm / 1064 nm wavelengths (phase index).
K(λ) is the dispersion equation of the standard air g(λ) is the dispersion equation of the water vapor
Edlen

## Slide 3
Arpent: a two-wavelength ADM operating at 780 nm and 1560 nm
3
7.4 kg
10.6 kg
9.8 kg
7.4 kg
0.4 kg
ADM and its Weather Station
Measurement head
Target, Al. hollow corner cube of 127 mm aperture
Target, Au. hollow corner cube of 127 mm aperture

## Slide 4
k=0
k=1
…
φ
One-way
Return
ADM
Target
Arpent: a two-wavelength ADM operating at 780 nm and 1560 nm
4
  At each wavelength:
  The beam is emitted continuously (CW laser) and is amplitude-modulated by a radio-frequency wave at 5 GHz .
  The OPL is determined by measuring the phase accumulated by the radio frequency wave fRF .
  This RF signal is propagated in the air by means of an optical beam of wavelength λ, i.e. at the speed of light c .
φ

## Slide 5
Conclusion from IWAA 2024 and Developments Since Then
Traceable to the International System of Units (SI) definition of the metre. This is ensured by knowledge of the RF wave locked to a Rubidium clock.
Instrumental uncertainty between 0.3 mm and 0.7 mm (k = 1) for SNRs from 60 dB to 50 dB. This was verified by field measurements.
Range up to 8 km.
12.3 km !
Tested on the CERN geodetic network in 2025.
5
Very satisfying result over kilometres, But not good enough over short distances up to 100 m.

## Slide 6
How to reduce the instrumental uncertainty ?
6

## Slide 7
we keep using ADM at λ = 1.5 µm u ~ 2 µm (k = 1)
difference in optical path lengths between the two wavelengths u = 6.7 µm (k = 1)
A = 48
The physical model (group index | 780 & 1560 nm)
New approach for the air-index compensation
7
we propose to use interferometry
at 532 nm and 1064 nm
u ~ 20 nm (k = 1)
A = 65
(phase index | 532 & 1064 nm)
fractional part
integer part
error of 10 °C ⇒ error in L1064 of 3 µm over 100 m
Frequency-doubled Nd:YAG laser
Notes: The Optical Path Length ("OPL") corresponds to a number of wavelength cycles accumulated by light over its path. This number of cycle can be splits into an integer part and a fractional part:

## Slide 8
Dispersion increase by 4.16 µm/m under standard atmospheric conditions (20 °C, 1013.25 hPa, 50%, 450 ppm)
4.16 µm/m
A hybrid ADM–IFM system… with a hard problem !
8
IFM: fractional part
IFM: integer part
ADM
65
65
The proposed approach is technically feasible
The hard problem !
u ADM ≪ 1 mm ⇒ u L 532 – L 1064 ≪ 4 nm
only true if the atmospheric parameters are well known
Dispersion can be calculated as :

## Slide 9
Quantification of the dispersion
IFM: fractional part
IFM: integer part
ADM
65
65
Variation in the atmospheric conditions | Change in dispersion
ΔT = 1°C | 13.4 nm/m
Δp = 1 hPa | -4.1 nm/m
ΔRH = 5% | 1.3 nm/m
ΔxCO2 = 100 ppm | 0.2 nm/m
9
Dispersion can be calculated as :

## Slide 10
Quantification of the dispersion
IFM: fractional part
IFM: integer part
ADM
65
65
50 m
± 5 °C ⇒ ± 3.4 µm
ADM Helps, but not solves the hard problem !
10
Dispersion can be calculated as :
We only take the temperature T into account. (1013.25 hPa, 50%, 450 ppm).
± 5 °C at 50 m ⇒ 26 possible candidates for

## Slide 11
In Summary
IFM: fractional part
IFM: integer part
ADM
65
65
11
Currently, the ADM can measure absolute distances,
while the IFM can track subsequent distance variations caused by changes in the air refractive index.
We are currently implementing a third IFM wavelength to resolve the ambiguity.
He-Ne laser at 633 nm
Experimental setup and Experimental results
The hard problem !

## Slide 12
Experimental Setup
12

## Slide 13
Distant Target
The hybrid ADM–IFM system
13
532 nm
1064 nm
1550 nm
Frequency-doubled Nd:YAG laser
ADM developed in-house
Optical Head
The interferometric polarized beams and the telemetric optical signal are transmitted to an optical head by optical fibers

## Slide 14
Top view with corrections and a laser beams added by AI to improve readability
14
IFM Ref. Target
532 nm
1064 nm
1550 nm
~ 29 mm
Dichroic Mirror
ADM Target
Dichroic Mirror
PBS
Double Round-Trip
Beam expander x5
Ø 1.1 mm at 532 nm | Ø 1.4 mm at 1064 nm
Ø 5.7 mm at 532 nm | Ø 6.9 mm at 1064 nm
Distant Target
Homodyne detection in quadrature at 532 nm
Homodyne detection in quadrature at 1064 nm
Focus on the optical head
λ / 4
λ / 4
Round-Trip

## Slide 15
Some photographs
15
Target, Al. hollow corner cube of 127 mm aperture
Weather Station
Optical Head, 35 cm × 40 cm, h = 10 cm

## Slide 16
Experimental Results
16
1. Comparison with a 3-m-long interferometric bench
2. Long-term measurement over 2 m
3. Long-term measurement over 12 m and 66 m

## Slide 17
ADM
Performance of the hybrid ADM–IFM system
17
Std. Dev. = NA, 1.2 µm and 2.7 µm
Std. Dev. = 1.0 µm, 1.2 µm and 1.1 µm
Experiment 1: ADM measurements only
The air-index-compensated distance

## Slide 18
Experimental Results
18
1. Comparison with a 3-m-long interferometric bench
2. Long-term measurement over 2 m
3. Long-term measurement over 12 m and 66 m

## Slide 19
2.05 m distance measurement over 5 days: IFM
19
weekend
IFM Measurement:
Dispersion (fractional part) evolves as expected.
Residuals show a standard deviation of σ = 5.0 nm.
T between 18.2 °C and 22.0 °C
P between 1002 hPa and 1019 hPa RH between 47 % and 58 %
Friday
Thursday
Monday
Tuesday
Wenesday
Setup:
The optical head and target were set up on an optical table. A weather station was placed at the midpoint of the baseline.
weather station

## Slide 20
2.05 m distance measurement over 5 days: ADM
20
Expected OPL:
A fixed mechanical distance of 2.05 m, evolving over time according to the air refractive index.
OPL Measurement:
The measured OPL evolves differently from the expected trend.
T between 18.2 °C and 22.0 °C
P between 1002 hPa and 1019 hPa RH between 47 % and 58 %
weekend
Friday
Thursday
Monday
Tuesday
Wenesday
Thermal expansion of the baseline ?
Thermally induced movement of the head or target ?

## Slide 21
2.05 m distance measurement over 5 days: ADM
21
T between 18.2 °C and 22.0 °C
P between 1002 hPa and 1019 hPa RH between 47 % and 58 %
weekend
Friday
Thursday
Tuesday
Wenesday
Expected OPL:
A fixed mechanical distance of 2.05 m, evolving over time according to the air refractive index.
OPL Measurement:
The measured OPL evolves differently from the expected trend.
Thermal expansion of the baseline ?
Thermally induced movement of the head or target ?

## Slide 22
Experimental Results
22
1. Comparison with a 3-m-long interferometric bench
2. Long-term measurement over 2 m
3. Long-term measurement over 12 m and 66 m

## Slide 23
weekend
Friday
Monday
Tuesday
Wenesday
Thursday
Friday
Thursday
Wed
A simple model of the thermal expansion of the concrete slab based on 3 parameters:
1. a time constant of 13 hours 45 minutes (thermal inertia of the material); 2. a thermal expansion coefficient of 2.7 µm/m/°C; 3. an offset that defines the distance of the slab at 0 °C, being ~900 µm below the initial air-index-compensated distance.
Air-index-compensated distance over 12 m and 8 days
23
T between 23.0 °C and 27.3 °C
P between 1006 hPa and 1022 hPa RH between 39 % and 54 %
Setup:
The optical head and target were set up on cast-iron pillars. Three weather stations were distributed along the baseline.
air-index-compensated distance
fitted model

## Slide 24
Air-index-compensated distance over 12 m and 8 days
The residuals after subtracting the fitted model:
distance variations between -20 µm et +30 µm
24
weekend
Friday
Monday
Tuesday
Wenesday
Thursday
Friday
Thursday
Wed
Standard deviation of 10.4 µm
T between 23.0 °C and 27.3 °C
P between 1006 hPa and 1022 hPa RH between 39 % and 54 %
The remaining variations are likely due to thermally induced movement of the optical head or the target ?
residuals
air-index-compensated distance
fitted model
=

## Slide 25
Air-index-compensated distance over 66 m and 14 hours
25
polynomial fit
residuals of 1.6 µm standard deviation
air-index-compensated distance
Estimation of the distance resolution over a long distance:

## Slide 26
Conclusion
26

## Slide 27
Conclusion
27
A new hybrid ADM-IFM system for air refractive index compensation An original approach that avoids determining the absolute values of the fringe orders and , requiring only the difference . The ADM, when combined with a weather station, provides a limited number of possible candidates for Next: solve the determination of ⇒ Add a third IFM wavelength at 633 nm
Range up to 66 m experimentally tested.
Promising Results: Short-distance measurements ⇒ uncertainty < 3 µm (k=1) for displacements up to 3 m. Long-distance measurements ⇒ resolution < 2 µm and reliable tracking of the dispersion by IFM over several days.
Remaining Challenges: Air-index-compensation appears effective, but distance variations due to mechanical instability / thermal expansion.

## Slide 28
Thank you for your attention !

## Slide 30
Bonus Slides
30

## Slide 31
IFM performance
31
Slope = 4.16 µm/m
STD = 43 nm for all points
STD = 21 nm if we select only points with IFM levels at 532 nm higher than 4x106
STD = 12 nm if we select only points with IFM levels at 532 nm higher than 6x106
IFM

## Slide 32
25.6 m
39.4 m
IFM: Lissajous figures
32
radius
The x-y plane, where each axis represents 10-6 of the ADC value from 0 to 223-1.
The observation time is 20 s.

## Slide 33
IFM: Lissajous figures
33
53.4 m + 1 cm
66.0 m
53.4 m
unfavourable distance and poor alignment
favourable distance and good alignment
Level at the lower limit for good performance
Phase accuracy degraded at this level

## Slide 34
2.4 µm
5.9 µm
3.6 µm
532 nm – 633 nm
48 candidates
532 nm – 1064 nm
22 candidates
633 nm – 1064 nm
30 candidates
IFM: Dispersion candidates
34
Curves obtained by a simulation