## Slide 1: SOLEIL II alignment: a metrological approach to meeting tolerances with confidence
Y. Delalande1, S. Ducourtieux1, A. Lestrade1, M. Sebdaoui1, B. Leluan1, C. Bourgoin1, Y. Denis1, D. Jourdain1, J. Guillory2, S. Durand3
1Synchrotron SOLEIL
2Laboratoire Commun de Métrologie LNE (Cnam/LCM - EA 2367)
3Laboratoire Geomatique et Foncier (Cnam/GeF - UR 4630)

## Slide 2: SOLEIL II project
CDR
(Conceptual Design Report)
TDR
(Technical Design Report)
STAGE 1:
CONSTRUCTION
STAGE 2:
TOWARDS FULL PERFORMANCE
2020
2021
2022
2023
2024
2025
2026
2027
2028
2029
2030
2031
2032
2033
2034
2035
Shutdown
Quadrupole SOLEIL/SOLEIL II
Shutdown and dismantling of the existing machine
Commissioning of the new booster and storage ring
Adaptation and restart of the beamlines
Information System Overhaul
Modernization of the support labs
R&D program on instrumentation
Sextupole SOLEIL/SOLEIL II
BPM vacuum chamber SOLEIL/SOLEIL II
Complete Modernization of the beamlines
Full performance of the accelerators
Modernization of the support labs
R&D program on instrumentation
Evolution towards innovative analysis techniques
SOLEIL
LINAC energy: 110 MeV
Booster: 0.1 - 2.75 GeV
SR circumference: 354.10 m
SR energy: 2.75 GeV
SR emittance: 3900 pm.rad
317 electromagnets
56 girders
29 beamlines
SOLEIL II
LINAC energy: 150 MeV
Booster: 0.15 - 2.75 GeV
SR circumference: 353.98 m
SR energy: 2.75 GeV
SR emittance: 90 pm.rad
1300 permanent/electro-magnets
86 girders
29 beamlines

## Slide 3: SOLEIL II project
CDR
(Conceptual Design Report)
TDR
(Technical Design Report)
STAGE 1:
CONSTRUCTION
STAGE 2:
TOWARDS FULL PERFORMANCE
2020
2021
2022
2023
2024
2025
2026
2027
2028
2029
2030
2031
2032
2033
2034
2035
Shutdown
Quadrupole SOLEIL/SOLEIL II
Shutdown and dismantling of the existing machine
Commissioning of the new booster and storage ring
Adaptation and restart of the beamlines
Information System Overhaul
Modernization of the support labs
R&D program on instrumentation
Sextupole SOLEIL/SOLEIL II
BPM vacuum chamber SOLEIL/SOLEIL II
Complete Modernization of the beamlines
Full performance of the accelerators
Modernization of the support labs
R&D program on instrumentation
Evolution towards innovative analysis techniques
SOLEIL
LINAC energy: 110 MeV
Booster: 0.1 - 2.75 GeV
SR circumference: 354.10 m
SR energy: 2.75 GeV
SR emittance: 3900 pm.rad
317 electromagnets
56 girders
29 beamlines
SOLEIL II
LINAC energy: 150 MeV
Booster: 0.15 - 2.75 GeV
SR circumference: 353.98 m
SR energy: 2.75 GeV
SR emittance: 90 pm.rad
1300 permanent/electro-magnets
86 girders
29 beamlines
New lattice very compact and very sensitive to misalignment => tight alignment tolerances

## Slide 4: Main alignment requirements for SOLEIL II
Install the new storage ring in the existing tunnel while keeping the source points of the 15 preserved beamlines at their current position (+/- 0.5 mm)
Meet the alignment tolerances set by the accelerator physicists to reduce critical discontinuities along the beam path and guarantee the expected electron beam performance
Errors simulated by accelerator physicists (normal distribution truncated at +/- 2σ)
Alignment tolerances defined as +/- 2σ of the simulated error, for consistency
s
x
y
Beam
Ultra-low-emittance lattice with strong, densely packed magnets: a few tens of µm of misalignment degrades the beam
Permanent-magnet quadrupoles cannot be tuned, hence tight longitudinal and circumference tolerances

## Slide 5: Main alignment requirements for SOLEIL II
Install the new storage ring in the existing tunnel while keeping the source points of the 15 preserved beamlines at their current position (+/- 0.5 mm)
Meet the alignment tolerances set by the accelerator physicists to reduce critical discontinuities along the beam path and guarantee the expected electron beam performance
Errors simulated by accelerator physicists (normal distribution truncated at +/- 2σ)
Alignment tolerances defined as +/- 2σ of the simulated error, for consistency
s
x
y
Beam
Ultra-low-emittance lattice with strong, densely packed magnets: a few tens of µm of misalignment degrades the beam
Permanent-magnet quadrupoles cannot be tuned, hence tight longitudinal and circumference tolerances
How can we guarantee that these requirements will be met with sufficient confidence?

## Slide 6: Metrological approach for SOLEIL II
A metrological approach is a quality assurance approach aimed to ensure the reliability of a single measurement or a measurement strategy. It involves analysing, quantifying and controlling uncertainty throughout the measurement process, to deliver results that are traceable to the SI as far as possible, documented, repeatable and compliant with the specified requirements.
At SOLEIL, this approach led to a few key questions:
  What exactly do the alignment errors simulated by the physicists represent, and do they reflect the actual alignment process?
  What exactly is a tolerance, and how do we decide whether it is met?
  Are the alignment tolerances achievable with our measurement chain?
This approach goes beyond pure alignment and requires close collaboration with accelerator physicists, mechanical engineers, and magnet experts.
Requirements
& discussions
Tolerance
definition
Modelling and
simulations
R&D
Calibrations
& Traceability
Measurements
& comparisons
Uncertainty
budget
Conformity
assessment
Optimisations
& improvements
Metrological approach
Future contribution

## Slide 7: Measurand definition
Measurand = quantity intended to be measured
  Requires a precise definition to avoid any ambiguity regarding what is to be measured
  Must match exactly the quantity to which the tolerance applies: physicists and aligners must be talking about the same thing
“Horizontal transverse misalignment of the magnetic centre of a quadrupole with respect to the girder axis , at the nominal temperature of 21 °C, when the machine is off.”
“Relative deviation between magnets”
Which magnets?
Between what?
What is it exactly?
In which direction?
Example of poorly defined measurand:
Example of a measurand defined at SOLEIL:
Under which conditions?
In which frame?

## Slide 8: Towards a more realistic alignment error model
Tolerances are derived from Monte Carlo beam simulations, in which alignment errors are applied in successive layers: magnets on their girder, then girders in the tunnel
Better and more realistic simulated beam performance, avoiding further tightening of alignment tolerances
Initial model: independent girders
New model: correlated girders
Each girder drawn independently:
Girder-to-girder discontinuity: σ = 37 µm, max = 149 µm
Each girder positioned relative to the previous one, as done with the laser tracker, with ring closure:
Girder-to-girder discontinuity: σ = 26 µm, max = 85 µm
Same input for both models: σ = 25 µm in translation, 10 µrad in rotation

## Slide 9: Conformity: are the tolerances met?
Tolerances
Probability
Magnet position
Theoretical position
Tolerance = interval of permissible values for a measurand to meet specified requirements.
At SOLEIL: +/- 2σ of the error simulated by the physicists

## Slide 10: Conformity: are the tolerances met?
Tolerances
Probability
Magnet position
Theoretical position
Measured position
Conform
Tolerance = interval of permissible values for a measurand to meet specified requirements.
At SOLEIL: +/- 2σ of the error simulated by the physicists

## Slide 11: Conformity: are the tolerances met?
Tolerances
Probability
Magnet position
Theoretical position
Measured position
Not conform
Tolerance = interval of permissible values for a measurand to meet specified requirements.
At SOLEIL: +/- 2σ of the error simulated by the physicists

## Slide 12: Conformity: are the tolerances met?
Tolerances
Probability
Magnet position
Theoretical position
Measured position
Conform at better than 99.9 %
Tolerance = interval of permissible values for a measurand to meet specified requirements.
At SOLEIL: +/- 2σ of the error simulated by the physicists
A measured value inside the tolerance not guarantee that the true value is: measurand uncertainty must be considered

## Slide 13: Conformity: are the tolerances met?
Tolerances
Probability
Magnet position
Theoretical position
Measured position
Conform at 70 %
Tolerance = interval of permissible values for a measurand to meet specified requirements.
At SOLEIL: +/- 2σ of the error simulated by the physicists
A measured value inside the tolerance not guarantee that the true value is: measurand uncertainty must be considered

## Slide 14: Conformity: are the tolerances met?
Tolerances
Probability
Magnet position
Theoretical position
Measured position
2u
2u
Acceptance interval
Conform at better than 95 %
Tolerance = interval of permissible values for a measurand to meet specified requirements.
At SOLEIL: +/- 2σ of the error simulated by the physicists
A measured value inside the tolerance not guarantee that the true value is: measurand uncertainty must be considered
Conformity criterion (JCGM 106): the tolerance is reduced by guard bands of 2u, defining an acceptance interval that ensures a conformity probability of at least 95 %.

## Slide 15: Are the tolerances achievable?
The uncertainty on the final measurand results from the whole alignment chain, not from a single measurement
Only a global uncertainty budget tells us, before construction, whether the tolerances can be met and where to improve
?
Survey
Laser tracker, adjustment software
Girder alignment in tunnel
Laser tracker, motorization, gauges
Magnet alignment on girder
Laser tracker, inclinometer, gauges
Fiducialisation
Keyence sensors, CMM, laser tracker
Magnetic characterisation
Stretched wire, laser tracker, inclinometer
U = ?
U = ?
U = ?
U = ?
U = ?

## Slide 16: Laser tracker uncertainty budget
The laser tracker is used at every alignment step: a dedicated uncertainty budget was built for this instrument alone
Five years of R&D, experimental characterisation, standardised protocols and periodic calibration have provided in-depth knowledge of the laser tracker
Example: Reflector centring error
1) Mechanical centre: reference sphere probed in the same nest
2) Optical centre: intersection of the three mirror planes probed
Centring error = vector from mechanical centre to optical centre
Validation: laser tracker vs CMM on 13 reflectors:
Identification
Quantification
Propagation

## Slide 17: Laser tracker uncertainty budget
14 points
Trueness is rarely assessed: without a stable, well-known reference shape, the main option is comparison with an instrument of equal or better uncertainty
3 laser trackers
Reference: calibrated CMM (LK Altera M, MPE 4.0 µm + L/350)
Set-up: 14 reflectors along a 3 m line, measured simultaneously by all instruments
Conditions: 21 +/- 0.1 °C, 30 measurement series for each instrument
Repeatability: σ < 1 µm for the CMM, σ < 2 µm for each laser tracker
Analysis: mean shapes compared in the 3 directions
s
x
y
Realistic laser tracker uncertainty: 4 µm (k = 1) over a 3 m girder, measured from 1.5 m

## Slide 18: Magnetic characterisation
Objective: define each magnet’s own reference frame (6 DOF) from its magnetic properties
Method: a stretched wire locates the magnetic axis and plane, and the magnet is shimmed to the bench frame
y
x
s
Vertical magnetic axis repeatability over 160 h
*Values for a magnet with a pole length of 150 mm and a pin spacing of 340 mm on the girder
Magnetic DOF (measured)
Mechanical DOF (set by realisation)
 |  |  |  |  | * | *
Bench alignment | ● |  |  | ● | ● | ●
Mechanical realisation | ● |  |  |  | ● | ●
Magnetic axis detection |  | ● | ● |  |  | 
Wire sag |  |  | ● |  |  | 
Thermal expansion | ● | ● | ● | ● | ● | ●
Magnet alignment uncertainty (k = 1) | 35 | 5 | 6 | 50 | 200 | 215
Bench alignment | ● |  |  | ● | ● | ●
Mechanical realisation | ● |  |  |  | ● | ●
Magnetic axis detection |  | ● | ● |  |  | 
Wire sag |  |  | ● |  |  | 
Thermal expansion | ● | ● | ● | ● | ● | ●
Magnet alignment uncertainty (k = 1) | 35 | 5 | 6 | 50 | 200 | 215
Identification
Quantification
Propagation
µm
µm
µm
µrad
µrad
µrad

## Slide 19: Fiducialisation
Objective: link the magnetic reference frame to mechanical references
Method: the wire position is transferred to the fiducials via optical sensors and a laser tracker
Keyence non-linearity over a 6 mm displacement
Identification
Quantification
Propagation
Magnetic axis
Keyence axis
Keyence measurements
CMM measurements
Laser tracker measurements
Fiducialisation
 |  |  |  |  | * | *
Laser tracker uncertainty |  | ● | ● |  |  | 
CMM uncertainty |  | ● | ● |  |  | 
Keyence non-linearity |  | ● | ● |  |  | 
Inclinometer uncertainty |  |  |  | ● |  | 
Reflector centring error |  | ● | ● |  |  | 
Magnet alignment uncertainty (k = 1) | - | 6 | 6 | 10 | - | -
Laser tracker uncertainty |  | ● | ● |  |  | 
CMM uncertainty |  | ● | ● |  |  | 
Keyence non-linearity |  | ● | ● |  |  | 
Inclinometer uncertainty |  |  |  | ● |  | 
Reflector centring error |  | ● | ● |  |  | 
Magnet alignment uncertainty (k = 1) | - | 6 | 6 | 10 | - | -
µm
µm
µm
µrad
µrad
µrad
*Values for a magnet with a pole length of 150 mm and a pin spacing of 340 mm on the girder
Magnetic DOF (measured)
Mechanical DOF (set by realisation)

## Slide 20: Magnet alignment on girder
Objective: position each magnet’s reference frame within the girder reference frame
Method: magnets are clamped against pins, measured with a laser tracker and inclinometer, and shimmed if needed
Clamping repeatability before and after a 10 µm shimming
Identification
Quantification
Propagation
 |  |  |  |  | * | *
Clamping repeatability | ● | ● | ● | ● | ● | ●
Pin position on the girder | ● |  |  |  |  | ●
Shimming uncertainty | ● | ● | ● | ● | ● | ●
Laser tracker uncertainty |  | ● | ● |  |  | 
Inclinometer uncertainty |  |  |  | ● |  | 
Magnet alignment uncertainty (k = 1) | 20 | 5 | 5 | 10 | 40 | 75
Clamping repeatability | ● | ● | ● | ● | ● | ●
Pin position on the girder | ● |  |  |  |  | ●
Shimming uncertainty | ● | ● | ● | ● | ● | ●
Laser tracker uncertainty |  | ● | ● |  |  | 
Inclinometer uncertainty |  |  |  | ● |  | 
Magnet alignment uncertainty (k = 1) | 20 | 5 | 5 | 10 | 40 | 75
µm
µm
µm
µrad
µrad
µrad
*Values for a magnet with a pole length of 150 mm and a pin spacing of 340 mm on the girder
Magnetic DOF (measured)
Mechanical DOF (set by realisation)
Clamping tool
Length gauge to check the X pin displacement
Laser tracker to check the position
10.2 µm
Before shimming
After shimming

## Slide 21: Girder alignment in the tunnel
Objective: set the girders at their nominal positions and minimise inter-girder discontinuities
Method: girders installed two by two, measured in the network frame, and finely alignment with motorisation
Monte Carlo simulation of the inter-girder discontinuity
Identification
Quantification
Propagation
Magnet position uncertainty | ● | ● | ● |  | ● | ●
Motorisation resolution |  | ● | ● | ● | ● | ●
Thermal expansion | ● | ● | ● | ● | ● | ●
Laser tracker uncertainty | ● | ● | ● | ● | ● | ●
Inclinometer uncertainty |  |  |  | ● |  | 
Girder alignment uncertainty (k = 1) | 67 | 14 | 14 | 10 | 4 | 4
Magnet position uncertainty | ● | ● | ● |  | ● | ●
Motorisation resolution |  | ● | ● | ● | ● | ●
Thermal expansion | ● | ● | ● | ● | ● | ●
Laser tracker uncertainty | ● | ● | ● | ● | ● | ●
Inclinometer uncertainty |  |  |  | ● |  | 
Girder alignment uncertainty (k = 1) | 67 | 14 | 14 | 10 | 4 | 4
µm
µm
µm
µrad
µrad
µrad
100 000 draws
Translations: discontinuity between consecutive girders
Rotations: girder orientation relative to nominal

## Slide 22: Survey: simulation model
The survey determines the final position of every component and decides on conformity
Survey uncertainty is hard to evaluate: thousands of observations combined through a network adjustment
A Monte Carlo simulation model was therefore developed
The model must be verified experimentally
Survey configuration
Stations, points, observations, environment
Laser tracker budget
Error distributions
Simulated observations
Randomly drawn errors
Network adjustment
Soleil6D
Uncertainty on indicators
Global shape, circumference, inter-girder, intra-girder, etc.
Monte Carlo (N drawn)

## Slide 23: Survey: repeatability campaign
7 surveys of the storage ring in 3 weeks (4 laser trackers, 462 stations, 35290 measurements)
Magnet alignment uncertainty (k = 1) | Measured | 2 µm | 1 µm | 1 µm | - | -
 | Simulated | 3 µm | 2 µm | 2 µm | - | -
Girder alignment uncertainty (k = 1) | Measured | 6 µm | 8 µm | 10 µm | 5 µrad | 3 µrad
 | Simulated | 6 µm | 9 µm | 11 µm | 6 µrad | 4 µrad
Circumference uncertainty (k = 1) | Measured | 167 µm |  |  |  | 
 | Simulated | 50 µm |  |  |  | 
Magnet alignment uncertainty (k = 1) | Measured | 2 µm | 1 µm | 1 µm | - | -
 | Simulated | 3 µm | 2 µm | 2 µm | - | -
Girder alignment uncertainty (k = 1) | Measured | 6 µm | 8 µm | 10 µm | 5 µrad | 3 µrad
 | Simulated | 6 µm | 9 µm | 11 µm | 6 µrad | 4 µrad
Circumference uncertainty (k = 1) | Measured | 167 µm |  |  |  | 
 | Simulated | 50 µm |  |  |  | 
Overall shape of the storage ring controlled to within +/- 80 µm during a survey
Model validated on local indicators, circumference gap likely due to machine movements

## Slide 24: Final uncertainty estimation
Step uncertainties combined over the whole chain
Expanded uncertainty (k = 2) compared with each tolerance: probability of conformity, assuming no systematic bias (centred distributions)
Conformity guaranteed (> 95 %) for most measurands
Issues with Rx and Ry orientations of magnets
Ongoing discussion with the accelerator physicists: angular tolerances proportional to magnet length
  Alignment : tolerances matched to what can be achieved
  Beam: tighter tolerances on long magnets (dipoles), which have the largest impact on beam dynamics
Current estimate, to be updated as the alignment process and the uncertainty budgets evolve

## Slide 25: Conclusion and perspectives
Conclusion:
For a process as complex as the alignment of a fourth-generation synchrotron, the metrological approach proved rich and useful:
  clear definition of measurands, tolerances and decision rules
  questioning of every step of the process
  characterisation and optimisation of instruments and measurement processes
Building a complete uncertainty budget across the whole chain gives a global view that no single step provides
Results are very encouraging and confirm that the SOLEIL II alignment tolerances can be met
Perspectives:
Characterisation of machine movements, to better account for them during alignment
Magnetic characterisation and fiducialisation benches under construction, installation planed for august of 2027
Tolerances can’t be met without a good control of uncertainties

## Slide 26: Acknowledgements
Many thanks to all my colleagues at SOLEIL and more specifically those in the Alignment and Metrology Group:
Sebastien DUCOURTIEUX
Alain LESTRADE (technic advisor 20%)
Mourad SEBDAOUI
Bruno LELUAN
Cedric BOURGOIN
Yann DENIS
David JOURDAIN (CDD)
Vincent GIGOT (trainee)
Special thanks to my thesis supervisor, Stéphane Durand, and my advisor, Joffray Guillory

## Slide 27: Discussion
Submitted to MST and under review
Y. Delalande et al., "Optimisation of the SOLEIL II accelerator alignment strategy through the implementation of a metrological approach", under review in Meas. Sci. Technol.