## Slide 1: SOLEIL II alignment: a metrological approach to meet tolerances with confidence
Y. Delalande1, S. Ducourtieux1, A. Lestrade1, M. Sebdaoui1, B. Leluan1, C. Bourgoin1, Y. Denis1, D. Jourdain1, J. Guillory2, S. Durand3
1Synchrotron SOLEIL
2Laboratoire Commun de Métrologie LNE
3Laboratoire Geomatique et Foncier

## Slide 2: Outline
SOLEIL II project and alignment requirements
The metrological approach and its three key questions
  Measurand definition and alignment error model
  Tolerances and conformity assessment
  Uncertainty budget: laser tracker, alignment steps, survey, final estimate
Conclusion and perspectives
An overview of the approach, not a detailed description of each step

## Slide 3: SOLEIL II project
CDR
(Conceptual Design Report)
TDR
(Technical Design Report)
STAGE 1:
CONSTRUCTION
STAGE 2:
TOWARDS FULL PERFORMANCE
2020
2021
2022
2023
2024
2025
2026
2027
2028
2029
2030
2031
2032
2033
2034
2035
Shutdown
Quadrupole SOLEIL/SOLEIL II
Shutdown and dismantling of the existing machine
Commissioning of the new booster and storage ring
Adaptation and restart of the beamlines
Information System Overhaul
Modernization of the support labs
R&D program on instrumentation
Sextupole SOLEIL/SOLEIL II
BPM vacuum chamber SOLEIL/SOLEIL II
Complete Modernization of the beamlines
Full performance of the accelerators
Modernization of the support labs
R&D program on instrumentation
Evolution towards innovative analysis techniques
SOLEIL
LINAC energy: 110 MeV
Booster: 0.1 - 2.75 GeV
SR circumference: 354.10 m
SR energy: 2.75 GeV
SR emittance: 3900 pm.rad
317 electromagnets
56 girders
29 beamlines
SOLEIL II
LINAC energy: 150 MeV
Booster: 0.15 - 2.75 GeV
SR circumference: 353.98 m
SR energy: 2.75 GeV
SR emittance: 90 pm.rad
1300 permanent/electro-magnets
86 girders
29 beamlines

## Slide 4: SOLEIL II project
CDR
(Conceptual Design Report)
TDR
(Technical Design Report)
STAGE 1:
CONSTRUCTION
STAGE 2:
TOWARDS FULL PERFORMANCE
2020
2021
2022
2023
2024
2025
2026
2027
2028
2029
2030
2031
2032
2033
2034
2035
Shutdown
Quadrupole SOLEIL/SOLEIL II
Shutdown and dismantling of the existing machine
Commissioning of the new booster and storage ring
Adaptation and restart of the beamlines
Information System Overhaul
Modernization of the support labs
R&D program on instrumentation
Sextupole SOLEIL/SOLEIL II
BPM vacuum chamber SOLEIL/SOLEIL II
Complete Modernization of the beamlines
Full performance of the accelerators
Modernization of the support labs
R&D program on instrumentation
Evolution towards innovative analysis techniques
SOLEIL
LINAC energy: 110 MeV
Booster: 0.1 - 2.75 GeV
SR circumference: 354.10 m
SR energy: 2.75 GeV
SR emittance: 3900 pm.rad
317 electromagnets
56 girders
29 beamlines
SOLEIL II
LINAC energy: 150 MeV
Booster: 0.15 - 2.75 GeV
SR circumference: 353.98 m
SR energy: 2.75 GeV
SR emittance: 90 pm.rad
1300 permanent/electro-magnets
86 girders
29 beamlines
Very compact lattice, highly sensitive to misalignment => tight alignment tolerances

## Slide 5: Main alignment requirements for SOLEIL II
Install the new storage ring in the existing tunnel while keeping the source points of the 15 preserved beamlines at their current position (+/- 0.5 mm)
Meet the alignment tolerances set by the accelerator physicists to reduce critical discontinuities along the beam path and guarantee the expected electron beam performance
Errors simulated by accelerator physicists (normal distribution truncated at 2σ)
Alignment tolerances defined as +/- 2σ of the simulated error, for consistency
s
x
y
Beam
Ultra-low-emittance lattice with strong, densely packed magnets: a few tens of µm of misalignment degrades the beam
Permanent-magnet quadrupoles cannot be tuned, hence tight longitudinal and circumference tolerances

## Slide 6: Main alignment requirements for SOLEIL II
Install the new storage ring in the existing tunnel while keeping the source points of the 15 preserved beamlines at their current position (+/- 0.5 mm)
Meet the alignment tolerances set by the accelerator physicists to reduce critical discontinuities along the beam path and guarantee the expected electron beam performance
Errors simulated by accelerator physicists (normal distribution truncated at 2σ)
Alignment tolerances defined as +/- 2σ of the simulated error, for consistency
s
x
y
Beam
Ultra-low-emittance lattice with strong, densely packed magnets: a few tens of µm of misalignment degrades the beam
Permanent-magnet quadrupoles cannot be tuned, hence tight longitudinal and circumference tolerances
How can we guarantee that these requirements will be met with sufficient confidence?

## Slide 7: Metrological approach for SOLEIL II
A metrological approach is a quality assurance approach aimed to ensuring the reliability of a single measurement or a measurement strategy. It involves analysing, quantifying and controlling uncertainty throughout the measurement process, to deliver results that are traceable to the SI as far as possible, documented, repeatable and compliant with the specified requirements.
At SOLEIL, this approach led to a few key questions:
  What exactly do the alignment errors simulated by the physicists represent, and do they reflect the actual alignment process?
  What exactly is a tolerance, and how do we decide whether it is met?
  Are the alignment tolerances achievable with our measurement chain?
Requirements
& discussions
Tolerance
definition
Modelling and
simulations
R&D
Calibrations
& Traceability
Measurements
& comparisons
Uncertainty
budget
Conformity
assessment
Optimisations
& improvements
Metrological approach
Future contribution

## Slide 8: Measurand definition
Measurand = quantity intended to be measured (VIM)
  Requires a precise definition to avoid any ambiguity regarding what is to be measured
  Must match exactly the quantity used in physicist simulations: physicists and aligners must be talking about the same thing
“Horizontal transverse misalignment of the magnetic centre of a quadrupole with respect to the girder axis , at the nominal temperature of 21 °C, when the machine is off.”
“Relative deviation between magnets”
Which magnets?
Between what?
What is it exactly?
In which direction?
Example of poorly defined measurand:
Example of a measurand defined at SOLEIL:
Under which conditions?
In which frame?

## Slide 9: Towards a more realistic alignment error model
Tolerances are derived from Monte Carlo beam simulations, in which alignment errors are applied in successive layers: magnets on their girder, then girders in the tunnel
More realistic and beneficial for both sides: better simulated beam performance for physicists and no further tightening of tolerances for aligners
Initial model: independent girders
New model: correlated girders
Each girder drawn independently:
Girder-to-girder discontinuity: σ = 37 µm, max = 149 µm
Each girder positioned relative to the previous one, as done with the laser tracker, with ring closure:
Girder-to-girder discontinuity: σ = 26 µm, max = 85 µm
Same input for both models: σ = 25 µm in translation, 10 µrad in rotation

## Slide 10: Conformity: are the tolerances met?
Tolerances
Probability
Magnet position
Theoretical position
Tolerance = interval of permissible values for a measurand to meet specified requirements.
At SOLEIL: +/- 2σ of the error simulated by the physicists

## Slide 11: Conformity: are the tolerances met?
Tolerances
Probability
Magnet position
Theoretical position
Measured position
Conforming
Tolerance = interval of permissible values for a measurand to meet specified requirements.
At SOLEIL: +/- 2σ of the error simulated by the physicists

## Slide 12: Conformity: are the tolerances met?
Tolerances
Probability
Magnet position
Theoretical position
Measured position
Non-conforming
Tolerance = interval of permissible values for a measurand to meet specified requirements.
At SOLEIL: +/- 2σ of the error simulated by the physicists

## Slide 13: Conformity: are the tolerances met?
Tolerances
Probability
Magnet position
Theoretical position
Measured position
Tolerance = interval of permissible values for a measurand to meet specified requirements.
At SOLEIL: +/- 2σ of the error simulated by the physicists
A measured value inside the tolerance does not guarantee the true value is: measurand uncertainty must be considered
Conforming with a probability better than 99.9 %

## Slide 14: Conformity: are the tolerances met?
Tolerances
Probability
Magnet position
Theoretical position
Measured position
Tolerance = interval of permissible values for a measurand to meet specified requirements.
At SOLEIL: +/- 2σ of the error simulated by the physicists
A measured value inside the tolerance does not guarantee the true value is: measurand uncertainty must be considered
Conforming with a probability around 70 %

## Slide 15: Conformity: are the tolerances met?
Tolerance = interval of permissible values for a measurand to meet specified requirements.
At SOLEIL: +/- 2σ of the error simulated by the physicists
A measured value inside the tolerance does not guarantee the true value is: measurand uncertainty must be considered
Conformity criterion (JCGM 106): the tolerance is reduced by guard bands of 2u, defining an acceptance interval that ensures a conformity probability of at least 95 %.
Tolerances
Probability
Magnet position
Theoretical position
Measured position
2u
2u
Acceptance interval
Conforming with a probability better than 95 %

## Slide 16: Are the tolerances achievable?
Global uncertainty budget of the alignment strategy to estimate the uncertainty of each measurand
The uncertainty on the final measurand results from the whole alignment chain, not from a single measurement
?
Survey
Laser tracker, adjustment software
Girder alignment in tunnel
Laser tracker, motorisation, gauges
Magnet alignment on girder
Laser tracker, inclinometer, gauges
Fiducialisation
Keyence sensors, CMM, laser tracker
Magnetic characterisation
Stretched wire, laser tracker, inclinometer
U = ?
U = ?
U = ?
U = ?
U = ?

## Slide 17: Laser tracker uncertainty budget
The laser tracker is used at every alignment step: a dedicated uncertainty budget was built for this instrument alone
Five years of R&D: more than 30 uncertainty contributions identified and quantified, through experimental characterisation, protocols and calibration
Example: Reflector centring error
1) Mechanical centre: reference sphere probed in the same nest
2) Optical centre: intersection of the three mirror planes probed
Centring error = vector from mechanical centre to optical centre
Validation: laser tracker vs CMM on 13 reflectors:
Identification
Quantification
Propagation

## Slide 18: Laser tracker uncertainty budget
14 points
Experimental verification of the laser tracker uncertainty budget on a typical girder-like shape, by comparison with a reference instrument (CMM)
3 laser trackers
Reference: calibrated CMM (LK Altera M, MPE 4.0 µm + L/350)
Set-up: 14 reflectors along a 3 m line, measured simultaneously by all instruments
Conditions: 21 +/- 0.1 °C, 30 measurement series for each instrument
Repeatability: σ < 1 µm for the CMM, σ < 2 µm for each laser tracker
Analysis: mean shapes compared in the 3 directions
s
x
y
Realistic laser tracker uncertainty estimated: 4 µm (k = 1) over a 3 m girder, measured from 1.5 m in a controlled environment

## Slide 19: Magnetic characterisation
Objective: define each magnet’s own reference frame (6 DOF) from its magnetic properties
Method: a stretched wire locates the magnetic axis and plane, and the magnet is shimmed to the bench frame
y
x
s
Vertical magnetic axis repeatability over 160 h
*Values for a magnet with a pole length of 150 mm and a pin spacing of 340 mm on the girder
Magnetic DOF (measured)
Mechanical DOF (set by realisation)
 |  |  |  |  | * | *
Bench alignment | ● |  |  | ● | ● | ●
Mechanical realisation | ● |  |  |  | ● | ●
Magnetic axis detection |  | ● | ● |  |  | 
Wire sag |  |  | ● |  |  | 
Thermal expansion | ● | ● | ● | ● | ● | ●
Magnet alignment uncertainty (k = 1) | 35 | 5 | 6 | 50 | 200 | 215
Bench alignment | ● |  |  | ● | ● | ●
Mechanical realisation | ● |  |  |  | ● | ●
Magnetic axis detection |  | ● | ● |  |  | 
Wire sag |  |  | ● |  |  | 
Thermal expansion | ● | ● | ● | ● | ● | ●
Magnet alignment uncertainty (k = 1) | 35 | 5 | 6 | 50 | 200 | 215
Identification
Quantification
Propagation
µm
µm
µm
µrad
µrad
µrad

## Slide 20: Fiducialisation
Objective: link the magnetic reference frame to mechanical references
Method: the wire position is transferred to the fiducials via optical sensors and a laser tracker
Keyence non-linearity over a 6 mm displacement
Identification
Quantification
Propagation
Magnetic axis
Keyence axis
Keyence measurements
CMM measurements
Laser tracker measurements
Fiducialisation
 |  |  |  |  | * | *
Laser tracker uncertainty |  | ● | ● |  |  | 
CMM uncertainty |  | ● | ● |  |  | 
Keyence non-linearity |  | ● | ● |  |  | 
Inclinometer uncertainty |  |  |  | ● |  | 
Reflector centring error |  | ● | ● |  |  | 
Magnet alignment uncertainty (k = 1) | - | 6 | 6 | 10 | - | -
Laser tracker uncertainty |  | ● | ● |  |  | 
CMM uncertainty |  | ● | ● |  |  | 
Keyence non-linearity |  | ● | ● |  |  | 
Inclinometer uncertainty |  |  |  | ● |  | 
Reflector centring error |  | ● | ● |  |  | 
Magnet alignment uncertainty (k = 1) | - | 6 | 6 | 10 | - | -
µm
µm
µm
µrad
µrad
µrad
*Values for a magnet with a pole length of 150 mm and a pin spacing of 340 mm on the girder
Magnetic DOF (measured)
Mechanical DOF (set by realisation)

## Slide 21: Magnet alignment on girder
Objective: position each magnet’s reference frame within the girder reference frame
Method: magnets are clamped against pins, measured with a laser tracker and inclinometer, and shimmed if needed
10 µm shimming with laser marking and cutting machine
Identification
Quantification
Propagation
 |  |  |  |  | * | *
Clamping repeatability | ● | ● | ● | ● | ● | ●
Pin position on the girder | ● |  |  |  |  | ●
Shimming uncertainty | ● | ● | ● | ● | ● | ●
Laser tracker uncertainty |  | ● | ● |  |  | 
Inclinometer uncertainty |  |  |  | ● |  | 
Magnet alignment uncertainty (k = 1) | 20 | 5 | 5 | 10 | 40 | 75
Clamping repeatability | ● | ● | ● | ● | ● | ●
Pin position on the girder | ● |  |  |  |  | ●
Shimming uncertainty | ● | ● | ● | ● | ● | ●
Laser tracker uncertainty |  | ● | ● |  |  | 
Inclinometer uncertainty |  |  |  | ● |  | 
Magnet alignment uncertainty (k = 1) | 20 | 5 | 5 | 10 | 40 | 75
µm
µm
µm
µrad
µrad
µrad
*Values for a magnet with a pole length of 150 mm and a pin spacing of 340 mm on the girder
Magnetic DOF (measured)
Mechanical DOF (set by realisation)
Clamping tool
Length gauge to check the X pin displacement
Laser tracker to check the position
10.2 µm
Before shimming
After shimming

## Slide 22: Girder alignment in the tunnel
Objective: set the girders at their nominal positions and minimise inter-girder discontinuities
Method: girders installed two by two, measured in the network frame, and finely aligned with motorisation
Monte Carlo simulation of the inter-girder discontinuity
Identification
Quantification
Propagation
Magnet position uncertainty | ● | ● | ● |  | ● | ●
Motorisation resolution |  | ● | ● | ● | ● | ●
Thermal expansion | ● | ● | ● | ● | ● | ●
Laser tracker uncertainty | ● | ● | ● | ● | ● | ●
Inclinometer uncertainty |  |  |  | ● |  | 
Girder alignment uncertainty (k = 1) | 67 | 14 | 14 | 10 | 4 | 4
Magnet position uncertainty | ● | ● | ● |  | ● | ●
Motorisation resolution |  | ● | ● | ● | ● | ●
Thermal expansion | ● | ● | ● | ● | ● | ●
Laser tracker uncertainty | ● | ● | ● | ● | ● | ●
Inclinometer uncertainty |  |  |  | ● |  | 
Girder alignment uncertainty (k = 1) | 67 | 14 | 14 | 10 | 4 | 4
µm
µm
µm
µrad
µrad
µrad
100 000 draws
Translations: discontinuity between consecutive girders
Rotations: girder orientation relative to nominal

## Slide 23: Survey: simulation model
The survey determines the final position of every component and decides on conformity
Survey uncertainty affects every measurand and is hard to evaluate
A Monte Carlo simulation model was therefore developed
The simulations must be verified experimentally
Survey configuration
Stations, points, observations, environment
Laser tracker budget
Error distributions
Simulated observations
Randomly drawn errors
Network adjustment
Soleil6D
Uncertainty on indicators
Global shape, circumference, inter-girder, intra-girder, etc.
Monte Carlo (N drawn)

## Slide 24: Survey: repeatability campaign
7 surveys of the storage ring in 3 weeks (4 laser trackers, 462 stations, 35290 measurements)
Magnet alignment uncertainty (k = 1) | Measured | 2 µm | 1 µm | 1 µm | - | -
 | Simulated | 3 µm | 2 µm | 2 µm | - | -
Girder alignment uncertainty (k = 1) | Measured | 6 µm | 8 µm | 10 µm | 5 µrad | 3 µrad
 | Simulated | 6 µm | 9 µm | 11 µm | 6 µrad | 4 µrad
Circumference uncertainty (k = 1) | Measured | 167 µm |  |  |  | 
 | Simulated | 50 µm |  |  |  | 
Magnet alignment uncertainty (k = 1) | Measured | 2 µm | 1 µm | 1 µm | - | -
 | Simulated | 3 µm | 2 µm | 2 µm | - | -
Girder alignment uncertainty (k = 1) | Measured | 6 µm | 8 µm | 10 µm | 5 µrad | 3 µrad
 | Simulated | 6 µm | 9 µm | 11 µm | 6 µrad | 4 µrad
Circumference uncertainty (k = 1) | Measured | 167 µm |  |  |  | 
 | Simulated | 50 µm |  |  |  | 
Overall shape of the storage ring controlled to within +/- 80 µm during a survey
Simulations validated on local indicators, circumference gap likely due to machine movements

## Slide 25: Final uncertainty estimation
Step uncertainties combined over the whole chain
Expanded uncertainty (k = 2) compared with each tolerance: probability of conformity, assuming no systematic bias (centred distributions)
Conformity guaranteed (> 95 %) for most measurands
  Transverse magnet alignment still to be improved
  Main issue: magnet pitch and yaw (Rx and Ry)
Ongoing discussion with the accelerator physicists: angular tolerances proportional to magnet length
  Alignment: tolerances matched to what can be achieved
  Beam: tighter tolerances on long magnets (dipoles), which have the largest impact on beam dynamics
Current estimate, to be updated as the alignment process and the uncertainty budgets evolve

## Slide 26: Conclusion and perspectives
Conclusion:
For a process as complex as the alignment of a fourth-generation light source, the metrological approach is essential: it aims to take everything into account and quantify it, to give us the best chance of conformity
It gave us a better understanding of what the physicists actually require:
  precise definition of measurands and a more realistic alignment error model
  clear decision rules for conformity, with guard bands
  a realistic estimation of the uncertainty budget across the whole alignment chain
A long process, never really finished: results are encouraging, critical cases are identified, and optimisations are under way to improve them
Perspectives:
Characterisation of machine movements, to better account for them during alignment
Magnetic characterisation and fiducialisation benches under construction, installation planned for august of 2027

## Slide 27: Acknowledgements
Many thanks to all my colleagues at SOLEIL and more specifically those in the Alignment and Metrology Group:
Sebastien DUCOURTIEUX
Alain LESTRADE (technic advisor 25%)
Mourad SEBDAOUI
Bruno LELUAN
Cedric BOURGOIN
Yann DENIS
David JOURDAIN (CDD)
Vincent GIGOT (trainee)
Special thanks to my thesis supervisor, Stéphane Durand, and my advisor, Joffray Guillory

## Slide 28: Poster and ongoing peer review paper
Submitted to MST and under review
Y. Delalande et al., "Optimisation of the SOLEIL II accelerator alignment strategy through the implementation of a metrological approach", under review in Meas. Sci. Technol.
Poster for IWAA 2026

## Slide 29: Thank you!

## Slide 30: Backup
Conformity of the horizontal transverse position Tx of the magnetic centres of the magnets on a girder
  Which magnets meet the tolerances and are conforming according to JCGM 106 (probability above 95 %) ?
Measured magnetic center and associated uncertainty (10 µm at k = 2)
GIRDER
Magnet
s
x

## Slide 31: Backup
Conformity of the horizontal transverse position Tx of the magnetic centres of the magnets on a girder
  Which magnets meet the tolerances and are conforming according to JCGM 106 (probability above 95 %) ?
Measured magnetic center and associated uncertainty (10 µm at k = 2)
GIRDER
Magnet
s
x
Theoretical axis of the beam

## Slide 32: Backup
Conformity of the horizontal transverse position Tx of the magnetic centres of the magnets on a girder
  Which magnets meet the tolerances and are conforming according to JCGM 106 (probability above 95 %) ?
Measured magnetic center and associated uncertainty (10 µm at k = 2)
GIRDER
Magnet
s
x
Theoretical axis of the beam
Specified tolerance for Tx (+/- 20 µm)
40 µm

## Slide 33: Backup
Conformity of the horizontal transverse position Tx of the magnetic centres of the magnets on a girder
  Which magnets meet the tolerances and are conforming according to JCGM 106 (probability above 95 %) ?
Measured magnetic center and associated uncertainty (10 µm at k = 2)
GIRDER
Magnet
s
x
Theoretical axis of the beam
Specified tolerance for Tx (+/- 20 µm)
Acceptance limit (+/- 10 µm)
40 µm
20 µm

## Slide 34: Backup
Conformity of the horizontal transverse position Tx of the magnetic centres of the magnets on a girder
  Which magnets meet the tolerances and are conforming according to JCGM 106 (probability above 95 %) ?
Measured magnetic center and associated uncertainty (10 µm at k = 2)
GIRDER
Magnet
s
x
40 µm
Theoretical axis of the beam
Specified tolerance for Tx (+/- 20 µm)
Acceptance limit (+/- 10 µm)
20 µm
Within the acceptance limit (not to be realigned)
Outside the tolerances (to be realigned)
Outside the acceptance limit (to be realigned)

## Slide 35: Backup
Conformity of the horizontal transverse position Tx of the magnetic centres of the magnets on a girder
  Which magnets meet the tolerances and are conforming according to JCGM 106 (probability above 95 %) ?
Measured magnetic center and associated uncertainty (10 µm at k = 2)
GIRDER
Magnet
s
x
40 µm
Theoretical axis of the beam
Specified tolerance for Tx (+/- 20 µm)
Acceptance limit (+/- 10 µm)
20 µm
Within the acceptance limit (not to be realigned)
Outside the tolerances (to be realigned)
Outside the acceptance limit (to be realigned)
The smaller the measurand uncertainty, the easier it is to decide on conformity

## Slide 36: Backup
Chart: T1 vs HLS
 | 0.010693 | 0.013718 | 0.04354 | 0.0066573 | -0.034398 | -0.062318 | -0.0048449 | -0.010549 | -0.027002 | 0.071836 | 0.0316 | 0.014396 | -0.051752 | -0.066423 | 0.013399 | 0.0039723 | 0.0023549 | 0.051219 | 0.025079 | 0.042193 | -0.010529 | -0.026116 | -0.011468 | -0.034545 | -0.0097588 | 0.039595 | -0.02511 | -0.0061721 | 0.0016584 | 0.020378 | 0.018687 | -0.051656 | -0.010349 | -0.050962 | 0.013593 | 0.023234 | 0.047024
Chart: T1 vs HLS
| 0 | 9199.9048567669961 | 19364.075602761077 | 31345.929049321552 | 41533.610882532157 | 53535.077655617977 | 63730.985072830874 | 67951.307075432938 | 72335.841474201501 | 87812.083284640888 | 92960.802969315147 | 97439.072816340951 | 107482.68207135194 | 119266.31487868738 | 129247.00259107331 | 140963.50561322144 | 150898.93597989457 | 155059.15538729349 | 173774.91316805076 | 179357.4463369695 | 183627.44854678508 | 193397.39108579117 | 204941.59831158322 | 214805.92689613576 | 226502.20797413166 | 236524.29022711393 | 240800.40602938455 | 244123.56814334373 | 259511.06078239967 | 265786.98927337886 | 270201.94271578465 | 280441.99964013416 | 292646.90255648602 | 303145.60290278506 | 315643.34050933633 | 326377.37593193649 | 330993.05436518899 | 334124.58579189033
T1 | 364.398 | 364.9135 | 364.9297 | 364.9033 | 364.8968 | 364.8324 | 364.6086 | 364.6889 | 364.6744 | 364.5387 | 364.9879 | 365.1161 | 364.9577 | 365.0386 | 365.0906 | 365.1081 | 365.1961 | 365.1736 | 365.1351 | 365.3221 | 365.4494 | 365.2787 | 365.263 | 364.8039 | 365.031 | 364.946 | 364.7789 | 364.8948 | 365.1833 | 365.1839 | 365.7613 | 365.6278 | 365.2396 | 365.0563 | 365.0489 | 364.8713 | 364.7966 | 364.9198
HLS | 697.7674 | 698.2753 | 698.2572 | 698.2632 | 698.2949 | 698.2561 | 697.974 | 698.0599 | 698.062 | 697.8293 | 698.32 | 698.4665 | 698.3775 | 698.4778 | 698.4546 | 698.4875 | 698.5825 | 698.5135 | 698.5112 | 698.684 | 698.8661 | 698.7155 | 698.6897 | 698.2567 | 698.4617 | 698.3287 | 698.2265 | 698.4883 | 698.6109 | 698.6025 | 699.1602 | 699.0255 | 698.7031 | 698.4738 | 698.5008 | 698.2528 | 698.1659 | 698.2636

## Slide 37: Backup
5 m x 5 m
granite
T5 building will be equipped with 4 benches and 2 laser trackers to support the characterization and fiducialization rate of the magnets:
  1 doublet/day/bench,
  2 quadrupoles, reverse bend /day/bench,
  1 dipole/2 days/bench)
Benches will be more compact and fully automated for measurements.
The air surrounding the benches will be controlled at a temperature of 21°C ± 0.25°C to improve the overall stability of the measurements, and with low airflow (< 0.15 m/s) to reduce air turbulence that perturbs laser tracker measurements.
Aluminum will be replaced by granite when possible:
Coefficient of thermal expansion 4 times lower than aluminum (goes from 24 µm/K/m to only 6.2 µm/K/m).
Granite will offer a much better thermal inertia.

## Slide 38: Backup
Chart: Prism position deviation between CMM and laser tracker measurements
| K1 | K2 | K3 | K4 | K5 | K6
S | 0.6 | -0.6 | 0.7 | -0.7 | -0.3 | 0.2
X | -0.1 | 0.1 | 1.2 | -1 | -0.8 | 0.7
Y | 0.3 | -0.6 | 0.4 | 0.5 | -0.9 | 0.2
Fiducialization of the Keyence sensor system under a CMM
Aim: detect the offsets between the sensors and the fixed prisms (the zero positions)
Measurement subcontracted to CETIM:
CMM: Leitz Infinity 2.4m x 1.2m x 1m.
5 repeated measurements for CMM.
100 repeated measurements for tracker
Proposed expanded uncertainty for our measurement: 2 µm (k=2).
Determination of the Keyence sensor zero with an uncertainty of 1 µm (σ).
CMM probe with 1 or 3 mm sapphire spheres
Deviation between all measurements = 1 µm (σ)